Vision-Based Micrometer Positioning of an XY Stepper Stage

Nguyễn Mạnh Bảo1, Trần Quang Nguyên2, Nguyễn Thanh Danh1, Lê Đức Anh1 and Trần Quang Trung1

1 Department of Solid State Physics, Faculty of Physics and Engineering Physics, University of Science, VNU-HCM, 227 Nguyen Van Cu Street, Cho Quan Ward, Ho Chi Minh City, Vi’

etnam

2 Faculty of Computer Engineering, University of Information Technology, VNU-HCM, Quarter 34, Linh Xuan Ward, Ho Chi Minh City, Vietnam

Abstract. XY stages are widely used in  optical alignment, microscopy systems which need to be moved with high precision. Lead-screw mechanisms driven by stepper motors are good for low-cost applications. However when they are used without any feedback their accuracy is often limited by things like backlash and static friction. This study focuses on making a low-cost XY positioning system that uses 0.5mm lead screw, two-phase stepper motors and a 25:1 reduction gearbox controlled by STM32F103C8T6 microcontroller combined with a camera to control the system at micrometer level.

In the experiments, the study evaluates key factors affecting positioning including bidirectional positioning error, backlash, closed-loop error, position time with three modes: open-loop stepping, additive backlash compensation, and look-and-move vision-based fine positioning.

The results showed that when the system is used in open loop the maximum error is 7.8 micrometers on X axis and 8.6 micrometers on Y axis. In additive backlash mode, the error when moving in both directions reduced down to about 4.3 micrometers and 4.9 micrometers, repeateadly. With image feedback and fine positioning control the error is reduced to about 2 micrometers, repeatability is approximately 1 to 1.5 micrometers, and typical positioning is 1.8 to 2 seconds. These results demonstrate that combining low-cost mechanical components and computer vision algorithms  can make micrometer-scale positioning more reliable. This is a foundation for making more advanced micromanipulation systems in the future.

Keywords: Micrometer positioning,XY stage, stepper motor, visual feedback, closed-loop control, backlash compensation.

# Introduction

Precise planar positioning is essential in optical alignment, microscopy sample handling, micro-assembly, and micromanipulation. Commercial precision stages utilizing piezoelectric actuators, linear motors, or high resolution encoders can achieve sub-micrometers or nanometer-level performance[1]. However, their high cost and complexity limit the accessibility for education and research. Therefore, stepper motor with lead screw stages are widely used for their low cost, simple structure and ease of control. Although, combining microstepping, gear reduction, and a fine-pitch lead screw, the theoretical displacement per microstep can be much smaller than one micrometer, practical positioning accuracy is lowered by backlash in the screw, nut, and gearbox causes lost motion when the direction changes. Thus, open-loop stepping is often inadequate for final micrometer-scale alignment.

To increasing the motion accuracy, visual servoing and image feedback are applied as effective non-contact measurement techniques [2-4]. Many machine vision methods have been applied to two-dimensional error calibration [11] and coplanar point to point positioning [12], still limit research the positioning efficency with a  low-cost lead-srew stepper combination to a software backlash compensation with a calibrated-reticle look-and-move fine loop.

In this paper, the study will focus on positioning performance on the low-cost lead-screw stepper stages combined with software compensation and calibrated image feedback. The hardware consists of 0.5mm pitch lead srew, a 25:1 reduction gearbox, two-phase stepper motor, an STM32F103C8T6 motion controller and an analog CCD camera. A microscope calibration recticle with 10 µm graduation scales on four sides and a central 4×4 array of sixteen square cells of 50 µm side length serves as the visual ruler for scale verification and position evaluation. Three control modes are evaluated on the same platform: open-loop stepping (M0), additive backlash compensation (M1), and look-and-move vision-based fine positioning (M2).

The remainder of the paper is organized as follows. Section 2 reviews related work on precision XY stages, backlash compensation, visual feedback, and vision-based stage metrology. Section 3 describes the hardware and methods. Section 4 presents the experimental procedure and results. Section 5 concludes the paper.

# Related Work

Precision position systems generating micro or nano scale motions are fundamental in medical manipulation, scanning probe microscopy, optical alignment, microscopy and micromanipulation. A wide variety range of XY has been studied, including piezoelectric flexure stages, voice-coil tables, impact-drive mechanisms,… For instances, Abdi[1] proposed an XY stages based on impact drive mechanism(IDM) driven by  piezoelectric actuators excited by saw-tooth voltage pulses. While IDM has compact structure, long stroke with micro-scale step generation through coarse-and-fine positioning strategies, they require specialized driving electronics and exhibit complex non-linear behavior[1]. In this paper, instead of a piezoelectric impact-drive stage, the system uses a two-phase stepper motor, gearbox, and lead screw. This offers the proposed approach easier to implement in laboratories environment.

However, the weakness of lead-screw stepper stages is that their actual motion deviates from the ideal kinemetic model due to many factors like backlash in the gearbox, screw, mechanical compliance and microstep nonlinearity. A common strategy is software backlash compensation that inserts additional pulse after a directional reversal. However, its efficency relies on the approximately repeatable backlash.

Therefore, visual servoing is a well-known method to guide mechanical motion using optical feedback[2]. Image-based visual servoing computes control inputs directly from image features and succefully applied to dynamic multi-DOF platform [3], demonstrating that visual measurements can effectively close the control loop on mechanical stages when Cartesian sensing is limited or costly. In this study, the fixed camera serves as a calibrated metrology sensor. By measuring relative displacements to a microscope reticle, the residual spatial errors and corrective commands are computed.

# System Description and Methods

## Hardware Configuration

The platform (Fig. 1) consists of  two-phase stepper motors coupled to  25:1 reduction gearbox and a 0.5mm/rev lead screw. Motion is driven by a TB6600 driver under control of an STM32F103C8T6 microcontroller in ¼ microstepping mode. Visual feedback is captured by an industrial analog CCD camera with the resolution of 640x480 pixels. Furthermore, a GUI is intergated to visual live image, target entry in micrometers, control-mode selection, position measurement, and experimental logging. For calibration and evaluation, the recticle which contains 10 µm graduation scales along the sides and a central 4 × 4 array of square cells with 50 µm side length serves as a visual ruler

Fig. 1. Complete XY positioning system

## Kinematic conversion

The nominal linear displacement per microstep L is estimated from the standard geared lead-screw model[7,8]:

𝐿 =𝑃/(𝑁∗𝑚∗𝑖)	(1)

where P is the lead-screw pitch, N is the full steps per revolution, m is the microstepping factor and i is the gear ratio. With P = 0.5mm/rev, N = 200, m = 4, and i = 25, the ideal displacement is:

L=500μm200*4*25=0.025 μm/microstep	(2)

Therefore, 40 microsteps correspond to 1 µm of linear travel. The pulse command Pulse required for a target displacement ∆x in μm becomes:

Pulse=40*  ∆x	(3)

## Vision measurement

### Point to point scale calibration

The pixel-to-micrometer scale is obtained by point-to-point measurement on the microscope calibration reticle. Two image points with the Euclidean pixel distance dPixel and a known physical distance dTrue in μm between two points, the scale factor s is:

s=dTruedPixel μm/pixel	 (4)

From repeated measurement(Fig. 2), the operational scale factor is approximately  1.7 μm/pixel.

Fig. 2. Point to point calibration on the microscope reticle in the host GUI

### Go-to-point localization by template matching and sub-pixel refinement

For go-to-point positioning, a small image region surrounding the selected destination is used as a reference template. In each frames, normalize cross-correclation locates the template peak at integer pixel resolution[10]. A classical three-point parabolic estimator is applied to the correlation peak. For neighboring intensity values I-1, I0, and I1, the sub-pixel offset ∆u is [13,14]

∆u= I-1- I12(I-1-2I0+I1)	(5)

Given the sub-pixel matched location (u,v)  and the commanded target location (ut, vt), the planar positioning error (x,y) is evaluated using the calibrated scale (4):

x=su-ut,  y=s(v-vt)	(6)

To suppress mechanical vibration and optical noise, final coordiantes are obtained by averaging mesurements across five consecutive settling frames.

Fig. 3. Go-to-point positioning selects a target and uses template matching with sub-pixel refinement

## Control modes

Three modes are implemented and evaluated on the same hardware platform, as summarized in Table 1.

Table 1. Evaluated control modes

| Mode | Description | Feedback | Backlash handling |
| --- | --- | --- | --- |
| M0 | Open-loop stepping | None | None |
| M1 | Open-loop stepping with additive backlash compensation | None | Extra pulses upon motion reversal |
| M2 | Vision-based fine positioning | Camera measurement | Additive compensation during coarse and fine moves |

### Mode M0: Open-loop stepping

The commanded displacement is converted directly to a pulse count. The microcontroller outputs the pulses without utilizing position feedback. The camera is used only for evaluating the post-motion position.

### Mode M1: Additive backlash compensation

The direction of each command is compared against the direction of the previous command. If a directional reversal is detected, the controller inserts additional pulses to take up the mechanical gap [9]. For an identified axial backlash b and an ideal microstep displacement L (μm/microstep) from (1), the required number of compensation pulses N is

N=roundbL	(7)

The compensation pulses are treated as hidden take-up pulses and do not change the external commanded coordinate.

### Mode M2: Vision-based fine positioning

Mode M2 operates via a look-and-move feedback based on the template-matching measurement in Section 3.3. The operator first selects the target image coordinate, storing the corresponding template. A coarse move is performed with the same backlash-compensated stepping as M1, using the calibrated scale (4) to convert the selected image target into a micrometer command. The camera then localizes the template again by matching and sub-pixel refinement (5), and computes the residual planar error (6)  sends a smaller correction command. The sequence is repeated until the residual positioning fall within a threshold.

Fig. 4.  M0,M1 flow chart (B), M2 flowchart (A)

# Experimental and Results

## Experimental conditions and evalutaion matrix

All experimentals are conducted under fixed illumination and optical focus, ensuring the reticle is kept inside the camera‘s field of view. The same vision pipeline evaluates the final position in all modes. The vision-referenced error e [2,5,6] is defined as:

e=pmeas- ptarget	(8)

Table 2. Experimental test matrix

| Test | Objective | Target Mode | Main Performance Metric |
| --- | --- | --- | --- |
| T1 | Open-loop accuracy | M0 | Maximum absolute error |
| T2 | Bidirectional hysteresis | M0 | Approach direction gap |
| T3 | Backlash idntification | M0 | Effective compensation |
| T4 | Compensated positioning | M1 | Maximum bidirectional error |
| T5 | Closed-loop positioning accuracy | M2 | Final residual error |
| T6 | Positioning time | M2 | Time to reach 2 μm residual error |
| T7 | Repeatability | M2 | Standard deviation of final error |

## Open-loop performance (M0)

In T1, target displacement of ±50, ±100, ±150, and ±200 µm are executed without compensation or feedback. Although the ideal displacement is only 0.025 µm per microstep, the measured open-loop error is much larger. The maximum absolute error is 7.8 µm on X and 8.6 µm on Y. The root mean square error (RMSE) is 4.0 µm (X) and 4.3 µm (Y); while the mean error is 0.7 µm on X and -0.5 µm on Y. These values indicate that the dominant contribution is not a large constant bias but direction-dependent mechanical loss.

In T2 examines direction effect by approaching the same target from the opposite sides. The final measured positions still differ measurably, with a hysteresis gap of approximately 6.1 µm on X and 5.8 µm on Y. This confirms that reversal behavior remains a source of error in the open-loop mode.

In T3 estimates the effective backlash through reverse-probe measurements. After preloading one direction, the stage is commanded to move back in 1 µm increments until measurable motion appears. The identified backlash is 5.9 µm on X and 5.4 µm on Y, and with the nominal displacement of 0.025 µm per microstep, X will need 236 compensation pulses and 216 for Y.

In conclusion, M0 is simple and fast, but its accuracy is still limited by  backlash, friction, and reversal-dependent motion loss.

## Additive backlash-compensated performance (M1)

Mode M1 uses the backlash values in T3 as hidden pulses whenever the direction reverses. In this way, the controller can reduce the dead zone effect without a closed loop system.

In  T4, the compensated mode is evaluated by repeating the bidirectional positioning tests with additive backlash correction enabled. The maximum absolute error is reduced to 4.3 µm on X and 4.9 µm on Y. RMSE drops to 2.2 µm and 2.4 µm, respectively, showing that most of the reversal-related positioning loss has been removed.

Mode M1 reduces succesfully the error band down to approximately 4-5 µm without requiring any extra sensor. However, a single constant compensation value still cannot present fully all friction values.

## Vision-based fine positioning performance (M2)

Mode M2 uses the compensated coarse motion M1 and then applies camera fine correction.

In T5, the final closed loop positioning accuracy is evaluated after the template matching and sub-pixel refinement. The maximum error is  2.1 µm on the X axis and 2.3 µm on the Y axis, while RMSE is approximately 1.0–1.1 µm. Mean error remains close to zero on both axes, indicating that the feedback loop effectively removes most systematic bias. These results are consistent with the calibrated imaging scale.

In T6, quantifies the time required for the feedback process. Positioning time is measured from the end of the coarse move to the first instant when the residual error is reduced to approximately 2 µm. Across 50 trials, the mean positioning time is 1.92 s, with a standard deviation of 0.16 s, a minimum of 1.63 s, and a maximum of 2.24 s. This is slower than pure open-loop motion because the controller must wait for settling and average multiple frames, but the final accuracy is significantly better.

In T7, by repeatedly returning to the same target from alternating approach directions evaluates repeability. The standard deviation of final error is 1.1 µm on X and 1.3 µm on Y. This remaining scatter is attributed to image noise, small illumination variation, mechanical residuals after the last correction.

Overall, M2 delivers the best positioning quality where the final residual error to about 2 µm and improves repeatability to about 1–1.5 µm. However, a longer positioning cycle and the need for a visible image feature is the limit.

## Comparision and Discussion

As shown in Table 3, the three modes exhibit a clear trade-off between simplicity, speed, and accuracy. M0 is the fastest but least accurate mode, M1 significantly reduces reversal-related error without extra sensing, and M2 delivers the smallest final residual error together with the best repeatability.

Table 3.  Performance Summary

| Mode | Main principle | Max error | RMSE | Repeatability | Typical time |
| --- | --- | --- | --- | --- | --- |
| M0 | Open-loop stepping | 7.8–8.6 µm | 4.0–4.3 µm | — | 0.3–0.6 s |
| M1 | Additive backlash compensation | 4.3–4.9 µm | 2.2–2.4 µm | — | Similar to M0 |
| M2 | Vision-based fine positioning | 2.1–2.3 µm | 1.0–1.1 µm | 1.1–1.3 µm | 1.8–2.0 s |

These results confirm that nominal microstep resolution does not directly determine the actual positioning accuracy of the stage. Moreover, additive backlash compensation removes a substantial portion of reversal-related error, but it cannot eliminate all residual uncertainty because a single constant compensation value cannot fully represent friction variation, compliance, or local screw imperfections. With the vision feedback, M2 is the most suitable mode for applications that prioritize final accuracy, while M0 and M1 remain attractive when simpler implementation or faster operation is more important.

The present results are obtained under controlled illumination using a calibration reticle with stable image features. In practical applications, performance may be affected by lighting variation, lower-contrast targets, or optical disturbances. Future work should therefore include validation with independent metrology, robustness testing under varying illumination, and faster convergence through improved imaging and control strategies.

# Conclusion

A low-cost XY positioning system based on a stepper motor, gearbox, lead screw, STM32F103 controller, TB6600 driver, and camera feedback was developed and evaluated. Although the ideal mechanical resolution is 0.025 µm per microstep, uncompensated open-loop positioning error reaches 7-9 µm because of backlash and friction. Additive backlash compensation reduces bidirectional error to approximately 4-5 µm. When image feedback and sub-pixel localization are added in a look-and-move loop, the final vision-referenced error about 2 µm, repeatability improves to about 1-1.5 µm, and typical positioning time is about 1.5–2 s.

The results demonstrate that a practical low-cost lead-screw XY stage combined with computer-vision-based correction offers a viable alternative to high-end precision stages. This approach is suitable for vision-guided applications such as microscope-based alignment and micromanipulation. Future work will focus on full-stroke calibration, robust vision processing, predictive correction, and workflow integration.

# References

1. Abdi, A.: Design and control of an XY precision positioning system using impact drive mechanism. Recent Patents on Mechanical Engineering 14 (2021). https://doi.org/10.2174/2212797614666210129150320

2. Pomares, J.: Visual servoing in robotics. Electronics 8(11), 1298 (2019). https://doi.org/10.3390/electronics8111298

3. Belmonte, Á., Ramón, J.L., Pomares, J., Garcia, G.J., Jara, C.A.: Optimal image-based guidance of mobile manipulators using direct visual servoing. Electronics 8(4), 374 (2019). https://doi.org/10.3390/electronics8040374

4. Ramón, J.L., Pomares, J., Felicetti, L.: Direct visual servoing and interaction control for a two-arms on-orbit servicing spacecraft. Acta Astronautica 192, 368–378 (2022). https://doi.org/10.1016/j.actaastro.2021.12.004

5. Hutchinson, S., Hager, G.D., Corke, P.I.: A tutorial on visual servo control. IEEE Transactions on Robotics and Automation 12(5), 651–670 (1996). https://doi.org/10.1109/70.538972

6. Chaumette, F., Hutchinson, S.: Visual servo control. I. Basic approaches. IEEE Robotics & Automation Magazine 13(4), 82–90 (2006). https://doi.org/10.1109/MRA.2006.250573

7. Acarnley, P.: Stepping Motors: A Guide to Theory and Practice, 4th edn. Institution of Engineering and Technology, London (2002).

8. Kenjo, T., Sugawara, A.: Stepping Motors and Their Microprocessor Controls, 2nd edn. Oxford University Press, Oxford (1994).

9. ISO 230-2:2014: Test code for machine tools — Part 2: Determination of accuracy and repeatability of positioning of numerically controlled axes. International Organization for Standardization, Geneva (2014).

10. Corke, P.: Robotics, Vision and Control: Fundamental Algorithms in MATLAB, 2nd edn. Springer, Cham (2017).

11. Wang, J., Zhao, W., Leach, R., Xu, L., Lu, W., Liu, X.: Positioning error calibration for two-dimensional precision stages via globally optimized image registration. Measurement 186, 110222 (2021). https://doi.org/10.1016/j.measurement.2021.110222

12. Ma, M.-Y., Huang, Y.-C., Wu, Y.-T.: A morphing point-to-point displacement control based on long short-term memory for a coplanar XXY stage. Sensors 23(4), 1938 (2023). https://doi.org/10.3390/s23041938

13. Willert, C.E., Gharib, M.: Digital particle image velocimetry. Experiments in Fluids 10, 181–193 (1991). https://doi.org/10.1007/BF00190388

14. Tian, Q., Huhns, M.N.: Algorithms for subpixel registration. Computer Vision, Graphics, and Image Processing 35(2), 220–233 (1986). https://doi.org/10.1016/0734-189X(86)90028-9